Application ID (as it appears in the application form / change notification form)

- [X, p.] in this document indicates a document to be named including page number – submitted for evidence. Grey text (for guidance) may be replaced/deleted.

- In case of a Change Notification, please only fill in the applicable sections. Please provide the latest full validation review / originating sterilization review project number (usually starting with 07xxxxxxxx)

- For most current version of Client Checklist please check Biological safety checklists | TÜV SÜD (tuvsud.com).

# Short Product Description relevant for Ethylene Oxide Sterilization

Note: Please replace italic text with respective information

Short description incl. picture of the device - in case of changes, as far as relevant

Description of the device as far as relevant for sterilization (pictures for clearer understanding): To be added Product schematic and / or photo of product, size, material, Intended Use / Intended Purpose according to IFU (inclusive total application duration, body contact, implantable, patient group), packaging description, picture Variants under assessment: To be added Product variants (e.g. Same product in different SBS, multiple products in same SBS) Description of the Sterile Barrier System specifications used at sterilization

Has this product previously been assessed by TÜV SÜD Product Service?

If yes, please provide 10-digit order no. usually starting with 071xxxxxxx, or equivalent traceable information

| Manufacturing facility and certification status of the applicable sterilization sites / facilities |  |
| --- | --- |
| Manufacturing site to be named | Please provide the applicable QMS Certificate 13485 of the used manufacturing site |
| Sterilization site to be named | Please provide the applicable QMS Certificate 13485 of the used sterilization site |

| External laboratories (if used for sterilization validation and routine and certification status of the laboratory) |  |
| --- | --- |
| Name of the laboratory | Please name the test done by the laboratory (e.g. microbiology BI testing, sterility testing, bioburden). Please provide the applicable Accreditation Certificate (e.g. ISO 17025 or GLP) |
| Name of the laboratory | Please name the test done by the laboratory (e.g. microbiology BI testing, sterility testing, bioburden). Please provide the applicable Accreditation Certificate (e.g. ISO 17025 or GLP) |

# Production related Information

## Equipment Specification

Note: Please replace italic text with respective information. Please add additional lines if required.

| Equipment including Identifier (e.g. int. ID/ serial number) | Site | Applicable cycle operated by the equipment for the device in question | Type of cycle | Usable chamber volume in m3 | All sensors / measurement devices (internal + external sensors, dataloggers for validation) are calibrated (statement sufficient) |
| --- | --- | --- | --- | --- | --- |
| e.g. Preconditioning (if applicable) | Inhouse or external source | Please name the cycle and version/ revision of cycle |  | 8 | ☐ Yes ☐ No |
| e.g. Sterilizer A10 | Inhouse or external source | Please name the cycle and version/ revision of cycle | e.g. (overpressure, underpressure) | 5,4 | ☐ Yes ☐ No |
| e.g. Venting Tunnel | Inhouse or external source | e.g. program #3 rev11 |  | 30 | ☐ Yes ☐ No |

## If Applicable: Cleaning of Product in Manufacturing before Sterilization

Note 1: Final cleaning is in scope of this section, which means cleaning before final packaging into sterile barrier system and sterilization. In consequence, intermediate cleaning steps are not in the scope of the sterilization assessment. Final cleaning could be performed via manual or automated cleaning processes.

Cleaning means reduction of manufacturing residues from the product i.e. grease, lubricants, particles, etc. An appropriate cleaning process is the basis for successful disinfection and sterilization, refer to section 2.3 Disinfection.

Note 2: Please replace italic text with respective information. Please add additional lines if required.

| General process information Final Cleaning |  |
| --- | --- |
| Are the products undergoing Final Cleaning before sterile barrier packaging and sterilization? | ☐ yes [X,p.y] ☐ no [X,p.y] if no, please mark N.A. in the table below for section 2.2. Description of Final Cleaning Process: Please add a short description of the final cleaning process, Final Cleaning process specification and procedure: Please add all references to the cleaning process procedure and specifications including the parameters of the cleaning process. Documented in [X,p.y] |
| Cleaning site | Please specify the site at which the Final Cleaning Process is located [X,p.y] ☐ N.A. |
| Cleaning process identifier | Please provide identification of the cleaning process in form of name and/or program number [X,p.y] ☐ N.A. |
| Cleaning equipment- if applicable | Please specify which cleaning equipment is used for the cleaning of product in scope with equipment name and device identifier [X,p.y] ☐ N.A. |
| Cleaning detergent- if applicable | Please specify the cleaning detergent [X,p.y], ☐ N.A. |
| Final Cleaning before Disinfection |  |
| Is the product disinfected after Final Cleaning? | ☐ yes [X,p.y] ☐ N.A. [X,p.y], please mark sections below in this chapter as N.A. |
| Are cleaning process validations/studies available providing evidence for efficient reduction of the manufacturing residuals? | ☐ yes [X,p.y] ☐ no ☐ N.A. Defined residuals to be reduced/removed: It shall be specified which residues could obstruct the subsequent disinfection process and are therefore intended to be reduced by final cleaning. Appropriate cleaning process is the basis for successful disinfection. Verification of successful reduction of defined manufacturing residues: Please add evidence for successful reduction of defined residues by validation, monitoring, and/or cleaning batch records. Documented in [X,p.y]. |

## If Applicable: Disinfection of Product in Manufacturing before Sterilization

Note 1: Disinfection means reduction of microbial load with a defined efficiency. Disinfection could be performed with chemical disinfection with disinfectant or physical disinfection like thermal treatment or rinsing with hot water

Note 2: Please replace italic text with respective information. Please add additional lines if required.

| General process information Disinfection ☐ N.A. [Section 2.3 can be omitted] |  |
| --- | --- |
| Disinfection process description and specification | Description of Disinfection Process: Please add a short description of the disinfection process Disinfection process specification and procedure: Please add reference to disinfection process procedure and specification including parameters of the disinfection process, specific load configurations, and supplies. Documented in [X,p.y]. |
| Disinfection site | Please specify the site at which disinfection process is located [X,p.y] |
| Disinfection process identifier | Please provide an identification of the disinfection process in form of name and/or program number [X,p.y] |
| Disinfectant- if applicable | Please specify the disinfectant [X,p.y] |
| Disinfection equipment- if applicable | Please specify the disinfection equipment in use for disinfecting product in scope with equipment name and device identifier [X,p.y] |
| Disinfection with defined efficiency for bioburden reduction |  |
| Are the products disinfected to achieve a defined efficiency for bioburden reduction, e.g. log reduction or A0 value? | ☐ yes [X,p.y] ☐ no |
| Are disinfection process validations/studies available providing evidence for efficient reduction of the microbiological contamination to a specified level? | ☐ yes [X,p.y] ☐ no Please provide documented evidence for efficient reduction of microbiological contamination of the disinfection process. Efficient reduction of bioburden could be shown by achieving specified log-reduction or A0 value. |

## If Applicable: Clean Room Control / Validation

Note: Please replace italic text with respective information. Please add additional lines if required.

This section is applicable to be filled in case of first evaluation of the clean room or in case of changes occurred to the clean room (e.g. new clean room, modification of the cleanroom and changes to the setup of the points listed in the table below).

| Is this the first evaluation of the clean room or have changes occurred to the cleanroom? (e.g. new clean room, modification of the cleanroom and changes to the setup of the points listed in the table below) | ☐ yes documented in [X,p.y] ☐ no, please mark N.A. in the table below for section 2.4. |
| --- | --- |
| Cleanroom | Please identify the cleanroom where the manufacturing takes place |
| Are action and alert limits set appropriately for the subsequent product bioburden in cleanroom processes? | ☐ yes documented in [X,p.y] ☐ N.A. ☐ no please justify: Acceptance criteria for “in operation” condition: Airborne particles [size]: particles/m3 Airborne microbiological contamination: cfu/m3 (and/or settle plates) Surface microbiological contamination: cfu/m2 Product bioburden: cfu (type – spores, fungi, anaerobe, bacteria) The bioburden shall be known to a degree to make decision on resistance |
| Monitoring points are defined for the above mentioned measurements | ☐ yes documented in [X,p.y] ☐ N.A. ☐ no please justify |
| Was IQ, OQ, PQ of the cleanroom successfully established? | ☐ yes documented in [X,p.y] ☐ N.A. ☐ no please justify: |
| Is all measuring equipment in a calibrated state? | ☐ yes documented in [X,p.y] ☐ N.A. ☐ no please justify: |
| Are utilities and media under surveillance | ☐ yes documented in [X,p.y] ☐ N.A. ☐ no please justify: Please specify what media and related acceptance criteria are defined. e.g. for water, compressed air… |
| Are environmental parameters defined | ☐ yes documented in [X,p.y] ☐ N.A. ☐ no please justify: Please specify where applicable: Temperature: Humidity: Pressure gradient: Air change rates: |

## Cycle Specification

Note: Please replace italic text with respective information for inhouse and outsourced processes.

Please add additional lines if required.

Please paste copy of the cycle specification used at routine sterilization:

Please paste here

| Is resterilization allowed? | ☐ Yes Please name the amount of re-sterilizations allowed ☐ No |
| --- | --- |
| Was product functionality verified after maximum amount of allowed sterilization cycles | ☐ Yes, documented in [x] ☐ No, please justify: |

## Basic Validation Development Data

Note: Please replace italic text with respective information.

| Validation Method | ☐ overkill half cycle approach (EN ISO 11135 Annex B 1.2.a) ☐ overkill cycle calculation approach (EN ISO 11135 Annex B 1.2b) ☐ BI and bioburden (EN ISO 11135 Annex A 1.3 b,c) ☐ Bioburden (EN ISO 11135 Annex A 1.3 a) ☐ Product adoption to an existing cycle – please fill box below Please assure that the adoption rationale and MPQ/PPQ validation data are submitted for the predicate device. This data is documented in [X] ☐ Other: Please specify and add rationale for not using a standardized method |
| --- | --- |
| Only for product adoption or cycle equivalence: ☐ N/A | Please provide the approved TÜV SÜD assessed predicate (EN ISO 11135, Annex D.12.5) project number: ______ Was a product adoption process followed and documented (e.g. based on AAMI TIR 28 or ISO 11135, D.12.5)? ☐ Yes, documented in [x] ☐ No Have aspects of effects on packaging integrity, product functionality and residuals been assessed? ☐ Yes, documented in [x] ☐ No Was a reduced validation performed based on chamber / process equivalence? ☐ Yes, rationale or data documented in [x] ☐ No |
| Revalidation criteria | ☐ By what events is a new validation is triggered - is documented in [X] Please provide the review interval of data, and interval of time till repeat MPQ / PPQ studies. When was the last MPQ, PPQ study executed? What are further criteria that to trigger a revalidation study (e.g. product changes…)? ☐ Validation protocol for the actual sterilization cycle including acceptance criteria - is documented in [X] |

## MPQ

Note: Please replace italic text with respective information for inhouse and outsourced processes. Please add additional lines if required.

Please paste copy of the cycle specification used at validation of sterilization MPQ:

Please paste here. Cycle record summaries are documented in [X]

| MPQ Processing - Microbial Performance Qualification |  |
| --- | --- |
| Biological indicators (if applicable) | ☐ Specification of Biological Indicators (BI) used is provided This data is documented in [X] Placement scheme and number of BIs or spore suspension, BI certificate ☐ Parametric release Placement of reference sensor(s) |
| BI results including culture conditions and time between end of cycle and BI testing | ☐ Sublethal cycle for BI resistance was performed – if no please justify: ☐ BI results including culture conditions and time between end of cycle and BI testing is documented in [x] |
| IPCD/EPCD (if applicable): | ☐ Drawings or pictures including place of inoculation position is provided in the submission including rational for the position and type taking into account: challenge to bioburden, physical conditions in load and diffusion pathways. This data is documented in [X] |
|  | ☐ Data on the relationship of the resistance between EPCD, IPCD, worst-case Products, natural Bioburden are provided. This data is documented in [X] |
| Bioburden | Please specify the bioburden level/limits e.g. in respect to bacteria, yeast/molds and anaerobic bacteria (were these investigations at least part of the initial assessment of bioburden) Please provide the bioburden trending data as a summary of the last year, if not available at least for the validation LOT. Please specify further for the environmental controlled manufacturing area the control limits for airborne bioburden and particles. |
| Is endotoxin testing applicable for the device under assessment? | ☐ no ☐ yes [X,p.y] Please specify the method and related results. The data is documented in [X]. (e.g. in case of direct contact to blood, CNS / CSF, eye or other systemic exposure) |

## PPQ Physical Performance Qualification

Note: Please replace italic text with respective information for inhouse and outsourced processes. Please add additional lines if required.

| Please specify the validation load configuration: | Please specify the load used during validation at PPQ and MPQ: Please consider min and max configuration in case of widely varying load configurations, scheme of total load), number of BIs, number of sensors (T and rH), scheme of position of BIs, total load volume, density, amount of adsorptive material (EO and water) (EN ISO 11135 9.4). The data is documented in [X] |
| --- | --- |
| Please specify which product was used in the load: | The product is the same as in section 1 ☐ yes [X,p.y] ☐ no – Please provide a description and justification |

Please paste copy of the cycle specification used at validation of sterilization PPQ:

Please paste here covering all of section 11135 9.5.4 The data is documented in [X]

Please assure that the following phases and process values and tolerances are part of the all over validation requirement.

Please be aware that the below parameters are not exhaustive to cover the cycle, cycle and load types, but are often omitted causing deficiencies and are thereof specifically requested.

| Phase | Acceptance criteria: values and tolerance | Results measured acc. 11135 9.5.4 | Comments if needed |
| --- | --- | --- | --- |
| Preconditioning: | Min/Max load temperature and humidity at end of preconditioning | To be added | To be added if any |
| Conditioning | Min/max load temperature and humidity during and end of conditioning | To be added | To be added if any |
| Sterilization/Exposure phase | Load temperature | To be added | To be added if any |
| The achievement of all cycle specification data and parameters above is verified during validation covering all of EN ISO 11135 9.5.4: | ☐ Yes, documented in [x] (best provided in a table showing setpoint/tolerances against measured data) ☐ No, please justify: |  | To be added if any |

# EO / ECH Residuals

Note: Please replace italic text with respective information for inhouse and outsourced processes.

Please add additional lines if required.

|  |  |  | Comments |
| --- | --- | --- | --- |
| Is a rationale provided for selection of representative sample(s) | ☐ Yes, documented in [X] | ☐ No, please justify: | To be added if any |
| Are the applicable Allowable limits for EO/ECH for the product in question considering the patient population provided | ☐ Yes, documented in [x] | ☐ No, please justify: | To be added if any |
| Results and test method provided | ☐ Yes, documented in [x] | ☐ No, please justify: | To be added if any |
| Has there been confirmatory testing of the degassing kinetics (at least 3 runs total) | ☐ Yes, documented in [x] | ☐ No, please justify: | To be added if any |

# Routine Processing

| Routine Processing |  |
| --- | --- |
| Please provide the EO-gas specification: | Please name the supplier of the EO gas and EO gas specification |
| Routine release | ☐ BI and physical parameters Placement scheme and number of BIs, BI certificate or spore suspension ☐ Parametric release Placement of reference sensor(s) |
|  | Release criteria documented in [x]: e.g. comparison of validation plot to routine plot, BI negative |
| Placement scheme and number of BIs, BI certificate provided? | ☐ Yes, documented in [x] ☐ No, please justify: |
| Routine cycle control and regulation is achieved by: | ☐ Chamber fixated sensors only ☐ Product integrated sensors (e.g. simulated product with integrated sensor) |
| Load configuration | ☐ Dedicated load Product configuration is fixed in number and location within chamber ☐ Mixed load different products allowed Please add information on min/max load variation, if applicable: |

| Release by client: |  |  |  |
| --- | --- | --- | --- |
|  | Date | Signature | Name |
|  |  |  | Name of Legal Manufacturer |

Note as to the signature’s relevance: If this document is officially signed, the provided rationales and data herein can be officially used by the reviewer. Otherwise, only the referenced documents can be used as evidence